## W3 - System Specific
all
System Specific Information |  |  |  |  | Some- please specify below
Tender Name | Envrionmental Decontamination
Lot Number | 1
Lot Name | Hydrogen Peroxide Vapour Disinfection Systems/Fogging Machines
If there are different responses required for different product categories please specify in additional columns
| Amity | BES Healthcare | Bioquell | Bioquell | Gama | Infinite | Infinite | Inivos | Innotec | Innotec | Sanital Med | UTB | UTB | UTB | UTB | UTB | Wassenburg Medical
Please provide model number / model name for each system you are submitting onto the framework. | Q82 - Airdecon 200 | Rhea Compact | Bioquell BQ-50 HC Suite (TD091-0000-230V) | Bioquell ProteQ (Sensor Version - TD092-0000-230V) | Scarlet-HPV | Viro Clean Easy | Viro Clean System | ProXcideTM | 2000.001- Nocospray | 3100.100 - Nocomax | RADS 1000 | Phileas 15 - (P15) | Phileas 25 - (P25) | Phileas 75 - (P75) | Phileas 250 - (P250) | Phileas 285 - (P285) | WASSENBURG GLOSAIR
System Specific Information | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response
Please Confirm brand name of your product | Airdecon 200 | Rhea Compact | Bioquell | Bioquell | SCARLETT by Clinell | VIRO CLEAN EASY | VIRO2 CLEAN SYSTEM |  | Oxy'Pharm | Oxy'Pharm | BioCair | PHILEAS | PHILEAS | PHILEAS | PHILEAS | PHILEAS | WASSENBURG GLOSAIR
Power supply (Voltage/Amps/Phase) | 230V, 3A, Single Phase | ~200/240V | Lectern HMI - 100-240V Single phase, 50/60 Hz, (0.4 A) Vaporiser - 230 V - Single phase, 50/60 Hz, (7.5A) Aeration Unit - 100-240 V Single phase,50/60 Hz, (3.1 A) Consumption: Lectern HMI - 0.006 kW Vaporiser - 1.8 kW Aeration Unit - 0.32 kW Supply: Installation category II | ProteQ Vaporiser Main Unit - 230 V single phase, 50/60Hz 7.5A, 1.8 kW ProteQ Lectern Unit - 230 V single phase, 50/60Hz 15A, 0.2 kW Optional Additional Aeration Unit - 100-240 V Single phase,50/60 Hz, (3.1 A) | 200-240VAC, 50/60Hz, 1205VA RHEA COMPACT; CUBEAIR 300VA single Phase | 1000 W | 2000 W | 230V/13A/Single Phase | 230v 13amp | 230v 13amp | 100-240VAC | NiMh Battery - 12V, 1.8amps | NiMh Battery and 240V/1.8amp/domestic - Transformer coverts usage to DC reducing power usage. Product is validated as Low Emissions and Low Power usage | 240V/1.8a/domestic - Transformer converts usage to DC reducing power usage. Product is validated as Low Emissions amnd Low Power usage | 240V/1.8a/ Transformer converts usage to DC reducing power usage. Product is validated as Low Emissions and Low Power usage | 240V/1.8a/ Transformer converts usage to DC reducing power usage. Product is validated as Low Emissions and Low Power usage | 2000W
Is your system designed to work on 110 converted or 220 | 220 | 220 | The system is available in 230V, 120V, or 100V versions | The system is available in 230V, 120V, or 100V versions | The system has been originally designed to operate on 220-240v electrical supply | 220 V/AC 50 Hz | 220 V/AC 50 Hz | 220-230V | 220 | 220 | 110 | Not applicable | 220 | 220 | 220 | 220 | 220
Heat output | 20oC | n/a | Same as power consumption listed above in kW | Same as power consumption listed above in kW | There is no data relating to heat output. There is no perceptible change in room ambient temperature post cycle. |  |  | Not yet tested, but known to be minimal | TBA | TBA | N/A | Minimal | Minimal | Minimal | Minimal | Minimal | +/- 37°C
Noise emission level | 60dB | n/a but assessed at <82dB | Noise at 1m: In room, at end of cycle 77.6dB(A), outside room 59.2dB(A) | Noise: Outside room over a cycle Laeq 64.2dB (A) | Aeration cycle noise (the majority of the cycle): 62dB | 83 Db | 83 Db | 85Db at 1m. As the system is designed to be used in sealed, unoccupied space, the exposure to the operator will be significantly lower than the stated output. | TBA | TBA | N/A | Below 30dBA | Below 30dBA | Below 30dBA | Below 35dBA | Below 35dBA | 100 +/-4dB
Machine size (height, width, length); | 104 cm x 40cm x 40cm | 397 x 340 x 561mm | Dimensions: Lectern HMI - 290 x 1040ax 220 mm (11.4 x 50ax 8.7 in) adjustable, minimum stated. Vaporiser Module - 450 x 560 x 460 mm (17.7 x 22 x 18.1 in). Aeration unit - 480 x 330 x 340 mm (18.9 x 13 x 13.4 in). Trolley - 700 x 1250 x 1350 mm (27.5x 49.2 x 53 in) | Dimensions: ProteQ Main Unit with Lectern - 672mm x 1430mm x 710mm (Height if no Lectern - 1421mm Optional Additional Aeration Unit - 480 x 330 x 340 mm | Injection unit: H 397 x L 561 x W 340 mm Aeration unit: H 700 x L 470 x W 590 mm | 43x32x37 cm | 43x32x37 cm | 1275mm H x 725mm W x 650mm D | 30x35x55cm | 103x59x69cm | 515 x 226 x 1040mm | 28x29x44 cms | 21x19x47 cms | 21x19x47 cms | 81x25x82 cms | 81x45x84 cms | H:974mm W:515mm D:566mm
Machine weight (kg). | 38kg | 8kg | Lectern HMI - 4.6Kg Vaporiser Module - 20Kg Aeration Unit - 10Kg | Weight: With lectern - 67kg Lectern alone - 9kg Optional Additional Aeration Unit - 10kg | Injection unit: 8kg Aeration unit: 22kg | 8.1 | 10.5 | 113kg | 7KG | 49KG | 38 kg | 11kgs including full 2 litre disinfectant chamber | 5kgs inlcuding full 500ml disinfectant chamber | 8kgs including full 2 litre disinfectant chamber | 25kgs including full 10 litre disinfectant canister | 48kgs including 2x full 10 litre disinfectant canisters | 33KG
Is your machine automated | Yes | Yes | Yes | Yes | Yes. Once the cycle parameters have been setup by the user the cycle will run fully automated (infection, contact time and aeration) | yes | yes | Yes | Yes | Yes | No | Yes | Yes | Yes | Yes | Yes | Yes
Is your product static or mobile during the cycle | Static | Static | Static | Static | Static | static | static | Static | Static | Static | Static | Static in use as required by UK BPR. But product is portable with carrying handles and less than the weight for a single person to carry. | Static in use as required by UK BPR. But product is portable with carrying handles and less than the weight for a single person to carry. | Static in use as required by UK BPR. But product is portable with carrying handles and less than the weight for a single person to carry. | Static in use as required by UK BPR. But product is mobile, on wheels with handles for propulsion.. | Static in use as required by UK BPR. But product is mobile, on wheels with handles for propulsion.. | Static
Number of emitters/carousels/towers that make up the system | 1 | 1 | 1 Vaporiser (Generator/Emitter) & 1 - 3 Aeration Units | 1 Vaporiser (Generator/Emitter) which houses an aeration unit. Additional aeration units can be added as required (2 can be housed inside of the vaporiser unit, a total of 6 can be connected.) | 1 | 1 | 2 | 1 | One | Four | 1 output for the dry-mist | Not applicable | Not applicable | Not applicable | Not applicable | Not applicable | 2
General Questions | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response
Please confirm the number of cycles which you would expect your machine to perform over its life (assuming its maintained in accordance with manufacturer's instructions & validated). | 18000 cycles | >6000 cycles | 10 year advised life span. Assuming 3 cycles per day, ~10,950 | 10 year advised life span. Assuming 3 cycles per day, ~10,950 | The equipment lifespan is expected to be at least 5 years. At an average usage of 20 cycles per week an approximate total of 5,200 cycles can be expected. | 550 working hour - 1.468.000 m3 treatment - 735 complete cycle of 45 minutes. | 550 wh - 2.936.000 m3 treatment - 735 complete cycles of 45 minutes. | 4000 - 5,000 | Number of cycles is not as relevant as expected hours of use as not all cycles are of the same length of time. We would expect approximately 1000 hours of use before the motor may need replacing - but this is just an estimate. We have numerous machines in the market which have been operating for many years, still with the original motor. | As Column B. | Approximately 80,000 Cycles assuming 20 cycles per day. | Phileas is designed to operate 24/7/365. The only limiting factors for cycles are - Down time to allow disinfectant to dwell and reach OEL for re-entry and availability of areas to be decontaminated | Phileas is designed to operate 24/7/365. The only limiting factors for cycles are - Down time to allow disinfectant to dwell and reach OEL for re-entry and availability of areas to be decontaminated | Phileas is designed to operate 24/7/365. The only limiting factors for cycles are - Down time to allow disinfectant to dwell and reach OEL for re-entry and availability of areas to be decontaminated | Phileas is designed to operate 24/7/365. The only limiting factors for cycles are - Down time to allow disinfectant to dwell and reach OEL for re-entry and availability of areas to be decontaminated | Phileas is designed to operate 24/7/365. The only limiting factors for cycles are - Down time to allow disinfectant to dwell and reach OEL for re-entry and availability of areas to be decontaminated | This is dependent on the size of the room.
Please provide the running cost per cycle for 100 cycles . | cost per cubic meter is £0.097, average room of 45m3 would cost £4.36. Therefore 100 cycles of this room would cost £436.00 | Depends on room size, protocol used and if adsorber (Cubair-OX) is used | Assuming a standard single room size of c.30m3 with average biocide dosing and the use of a Chemical Indicator in each room = c.£1,650 | Assuming a standard single room size of c.30m3 with average biocide dosing and the use of a Chemical Indicator in each room = c.£1,650 | Assuming a cost of £96 per bottle (2.5 litres of biocide): Assuming an infection rate of 6ml/m3 Assuming a room volume of 50m3 Cost per cycle/room: £11.52 Cost per 100 cycles/rooms: £1,152 |  |  | £20 per cycle | Running cost is determined by choice of chemical, concentration and size of room to be treated. It is unlikely 100 cycles would all be the same. Typically - a 50m3 room could be treated from £2.49 to £11.56 depending on the above. Hence, we would estimate the average cost for 100 treatments as £702.50 for a 50m3 room. | Running cost is determined by choice of chemical, concentration and size of room to be treated. It is unlikely 100 cycles would all be the same. Typically - a 500M3 room (note significant change in room volume to column A) could be treated from £15.91 to £80.62. depending on the above. | As per Apprendix 2 of the Tender, the cost for 100 cycles of a 57m3 Room would cost a total of £1,102 Excluding VAT. | Phileas 15 is ideal for areas with restricted Mains power, large Toilets, Bathrooms, Ambulances etc. with an operating range of 5-180 cubic metres. For 100 cycles in 40 cubic metre single rooms. Total cost £1776.00 | Phileas 25 is optimal for small rooms - up to 25 cubic metres and locations such as toilets and bathrooms. For 100 cycles up to 25 cubic metres each - £1110.00 | Phileas 75 is ideal for rooms/spaces from 5 - 180 cubic metres such as Rooms, Bays, Theatres, Treatment areas etc. For 100 cycles in 40 cubic metre single rooms. Total cost £1776.00 | Phileas 250 is ideal for larger volume spaces up from 50 to 650 cubic metres. Large bays, Therapy units, endoscopy suites etc. For 100 cycles in 200 cubic metre bays/spaces, Total cost £8880.00 | Phileas 285 is ideal for larger volume spaces up from 50 to 1100 cubic metres. Large bays, Therapy units, endoscopy suites, ITUs etc. For 100 cycles in 200 cubic metre bays/spaces, Total cost £8880.00. | This is dependent on the size of the room.
Please advise the cycle time of your product and for what room size (width/height/length) | 4m x 4m x3m = 169 minutes* | If used on 6ml/m3 in a room of 60m3 with a 30 min contact time = 40 minutes plus H2O2 extraction with the Cubair-OX (approx 1hr 10 mins) | Variable depending on individual enclosure to be decontaminated. Estimated cycle time for a 30m3 enclosure (4m x 3m x 2.5m) to be decontaminated to a 6 log sporicidal reduction level of kill is c.2 hours | Variable depending on individual enclosure to be decontaminated. Estimated cycle time for a 30m3 enclosure (4m x 3m x 2.5m) to be decontaminated to a 6 log sporicidal reduction level of kill is c.2 hours | An injection of 6ml/m³ is NF T72-281 compliant with a 6.95 MRSA log reduction in 70 minutes in a 30 m³ room | 10 m3 in 14 sec - 50 m3 in 1m 8 sec - 500 m3 in 11 m 16 sec - 1000 m3 in 22m 30 sec - 2000 m3 in 45 m | 10m3 in 7 sec - 50 m3 in 34 sec - 500 m3 in 5 m 38 sec - 1000 m3 in 11 m 15 sec - 2000 m3 in 22 m 30 sec - 4000 m3 in 45 min | Approximately 2 hours per cycle. Can treat rooms up to 170 cubic metres (10m W x 2.8m H x 6m L) | Cycle time is detremined by choice of chemical, concentration and size of room to be treated. A 50m3 room can be treated between 3 and 9 minutes depending on choice of chemical and concentration (ml/m3) selected. Accordingly, a 100m3 room can be treated between 6 and 18 minutes. | Cycle time is detremined by choice of chemical, concentration and size of room to be treated. A 500m3 room can be treated between 5 and 15 minutes depending on choice of chemical and concentration (ml/m3) selected. Accordingly, a 1000m3 room can be treated between 10 and 30 minutes. | 30 minute cycle for a 57m3 | For a 40 cubic metre room with a ceiling height up to 3 metres. Cycle time is - 55 minutes. Flow Rate is 12ml/hr and Max. room size capability - 180 cubic metres | For a 25 cubic metre room with a ceiling height of 3metres. Cycle time is - 65 minutes. Flow rate 500ml/hr and Max room size capability - 25 cubic metres | For a 40 cubic metre room with a ceiling height up to 3 metres. Cycle time is - 55 minutes. Flow Rate is 1200ml/hr and Max. room size capability - 180 cubic metres | For a 200 cubic metre 8- bedded Bay with height up to 4 metres. Cycle time is - 90 minutes. Flow rate is 3000ml/hr and Max room size capability 650 cubic metres | For a 200 cubic metre 8- bedded Bay with height up to 4 metres. Cycle time is - 85 minutes. For a 40 cubic metre room. Cycle time is 42 minutes. Flow rate is 4200ml/hr and Max room size capability 1100 cubic metres | Maximal treatment at once : 400m3 with 5ml/m3 according to test report norm EN17272 : 3h in total (1h diffusion + 2h contact time)
Is the cycle time for a sporicidal claim or a bactercidal claim | Sporcidal | Cycle time is individually tested for all pathogen groups | Sporicidal | Sporicidal | Please see Efficacy data attached for full details | YES; respecting 30 min contact time after end of the cycle | YES; respecting 30 min contact time after end of the cycle | Sporicidal and bactericidal | Cycle time is determined by choice of chemical - the lower cycle times are when operating a Preventative process, the higher cycle times are when operating a Curative process. | As Column B. | N/A | Both | Both | Both | Both | Both | Both
Does the emitter need to be repositioned at all during a cycle | No | No | No | No | No | NO | NO | No | No - once set up in accordance with recommendations, the machine requires no further attention until the cycle is complete. | As Column B. | N/A | Not relevant to this process | Not relevant to this process | Not relevant to this process | Not relevant to this process | Not relevant to this process | No
How does your machine record each cycle and reason for failure | Our unit has an on- board computer with built in track and trace system that tracks the user, cycle, and any failures. | Yes, up to 500 cycles | USB drive - provides alarm codes if cycle failure | Print out + real-time display. Provides alarm codes if cycle failure | Yes, data for each cycle is automatically collected in the internal memory and can be downloaded on USB and printed. | Complete record available for download on USB device and through free software | Complete record available for download on USB device and through free software | The system records every stage of every cycle, as well as the process information, such as room location, operator name, time of process etc. This produces a report which is saved to the cloud-based server, and emailed to a designated customer inbox. Further detailed information is stored on the server includind detailed process parameters, time for each stage of the cycle etc. Any failed cycles produce reports which clearly show how far the cycle proceeded. The information on the server can then be used to remotely assess the cause of the failure. | Machine is fitted as standard with a tracking system. This records the last 125 diffusions undertaken by the device (the 126th replaces the 1st and so on). This tells the operator the relevant information in terms of length of diffusion and chemical consumption. Validation of a successful diffusion is determined by use of Chemical Indicators (Nocotest). Should the CI's fail, there is a routine check list to be carried out, following which, if there is no obvious reason for failure we recomend the machine to be attended to by a trained technition. | As Column B. | N/A | The onboard software reports successful and failed cycle to the screen on the Phileas unit and to the MyPhileas App on a hand-held \tablet. The App/Table maintain the record in a down;aodable .pdf file for record-keeping purposes. | The onboard software reports successsful and failed cycle to the screen on the Phileas unit and to the MyPhileas App on a hand-held \tablet. The App/Table maintain the record in a down;aodable .pdf file for record-keeping purposes. | The onboard software reports successful and failed cycle to the screen on the Phileas unit and to the MyPhileas App on a hand-held \tablet. The App/Table maintain the record in a down;aodable .pdf file for record-keeping purposes. | The onboard software reports successful and failed cycle to the screen on the Phileas unit and to the MyPhileas App on a hand-held \tablet. The App/Table maintain the record in a down;aodable .pdf file for record-keeping purposes. | The onboard software reports successful and failed cycle to the screen on the Phileas unit and to the MyPhileas App on a hand-held \tablet. The App/Table maintain the record in a down;aodable .pdf file for record-keeping purposes. | Yes
Please confirm what certification your machine has been tested against | Electrical equipment for measurement, control, and laboratory use Low Voltage Directive (2006/95/EC) EMC Directive (2004/108/EC) | NF-T 72 201:2014 and EN:17272 2020 | System is CE & UK CA marked to applicable EMC and Low voltage regulations. In relation to standards, Bioquell HPV-AQ has been tested in combination with a BQ vapourisation module to EN 17272 and NFT 72-281 in accordance with EU & GB BPR. Bioquell has a European Biocidal Product Committee Union Authorisation approval opinion for its Bioquell HPV-AQ and vapor application system. | System is CE & UK CA marked to applicable EMC and Low voltage regulations. In relation to standards, Bioquell HPV-AQ has been tested in combination with a BQ vapourisation module to EN 17272 and NFT 72-281 in accordance with EU & GB BPR. Bioquell has a European Biocidal Product Committee Union Authorisation approval opinion for its Bioquell HPV-AQ and vapor application system. | NF-T 72-281: 2014 and EN17272:2020 | En13727, EN13624, EN 14348, AFNOR NF T -72 190, EN14476 | En13727, EN13624, EN 14348, AFNOR NF T -72 190, EN14476 | • BS EN12100:2010 • BS EN ISO 14971:2012 • BS EN 61010-1:2010 • BS EN 61010-2-201:2013 • BS EN 60601-1-2:2015 • BS EN 61326-1:2013 | NF T 72-281 & NF EN 17272 | As Column B. | EN 17272, EN13697:2015, EN1276: 2019, EN16777: 2018, EN14476:2013, EN13704: 2018 | AFNOR: NF T72-281 and BS:EN17272:2020 | AFNOR: NF T72-281 and BS:EN17272:2020 | AFNOR: NF T72-281 and BS:EN17272:2020 | AFNOR: NF T72-281 and BS:EN17272:2020 | AFNOR: NF T72-281 and BS:EN17272:2020 | IS0 13485, NFT 72281, EN17272
Please confirm date of certification | 2013-11-27 | 2014 | CE & UK CA mark = 12 Oct 2020. NFT 72-281 = 17 Feb 2016. EN 17272 = 15 June 2021. ECHA BPC approval opinion = 13 Oct 2021 | CE & UK CA mark = 12 Oct 2020. NFT 72-281 = 17 Feb 2016. EN 17272 = 15 June 2021. ECHA BPC approval opinion = 13 Oct 2021 | NF (2019) and the latest for the EN (2021) | see attached Supraspor technical report | see attached Supraspor technical report | 2018 | NF-T 72-281 - May 2015. NF EN 17272 - Nocolyse Neutral 4 June 2021 & Nocolyse One Shot 6 July 2021. | As Column B. | During 2020, 2021, 2022 | AFNOR 10 Jan. 2018 BS:EN17272 28 Apr 2021 | AFNOR 10 Jan. 2018 BS:EN17272 28 Apr 2021 | AFNOR 10 Jan. 2018 BS:EN17272 28 Apr 2021 | AFNOR 10 Jan. 2018 BS:EN17272 28 Apr 2021 | AFNOR 10 Jan. 2018 BS:EN17272 28 Apr 2021
Please confirm what you have tested your machine against from the list below: |  |  | Bioquell HPV-AQ + Bioquell vapourisation module | Bioquell HPV-AQ + Bioquell vapourisation module
Bactercidal activity | Tested - NFT 72 281 - log 7 reduction | Yes | Yes | Yes | Yes | yes | yes | Yes | Yes | As Column B. | yes | YES | YES | YES | YES | YES | Yes
Yeasticidal activity | Tested - NFT 72 281 - log 7 reduction | Yes | Yes | Yes | Yes | yes | yes | - | Yes | As Column B. | yes | YES | YES | YES | YES | YES | Yes
Sporicidal activity | Tested - NFT 72 281 - log 6 reduction | Yes | Yes | Yes | Yes | yes | yes | Yes | Yes | As Column B. | yes | YES | YES | YES | YES | YES | Yes
Mycobacterial activity | Tested - NFT 72 281 - log 6 reduction | Yes | Yes | Yes | Yes | yes | yes | - | Yes | As Column B. | yes | YES | YES | YES | YES | YES | Yes
Virucidal activity | Tested - EN 14476 - log 4 reduction | Yes | Yes | Yes | Yes | yes | yes | Yes | Yes | As Column B. | yes | YES | YES | YES | YES | YES | Yes
Fungicidal activity | Tested - NFT 72 281 - log 7 reduction | Yes | Yes | Yes | Yes | yes | yes | - | Yes | As Column B. | yes | YES | YES | YES | YES | YES | Yes
Please advise your log reduction rate on all of the above | Please see above | From 4 to 6 depending on protocol used | Bacteria >5log, Yeast >4log, Spores >4log, Mycobacteria >4log, Virucidal >4log, Fungicidal >4log | Bacteria >5log, Yeast >4log, Spores >4log, Mycobacteria >4log, Virucidal >4log, Fungicidal >4log | Please see Efficacy data attached for full details | see attached Supraspor technical report | see attached Supraspor technical report | >6log | In order to sell their equipment in EU/GB, ALL manufacturers of HPV systems are required to test their equipment against either standard NF T 722-281, NF EN 17272, or both. Oxy'Pharm Nocotech systems have both these accreditations. They each state the following pass rates: • Bactericidal activity – Log reduction ≥ 5 • Yeasticidal activity – Log reduction ≥ 4 • Sporicidal activity – Log reduction ≥ 3 • Mycobactericidal activity – Log reduction ≥ 4 • Virucidal activity – Log reduction ≥ 4 • Fungicidal activity – Log reduction ≥ 4 Their devices in many cases will exceed these benchmarks, but the test simply states ‘pass or fail’. Oxy’Pharm equipment has been subjected to numerous studies illustrating up to a Log 7 reduction, which greatly exceeds the requirements of the NF T norms. Details of this testing are available upon request. Furthermore, in March 2022, University of Cambridge published a Scientific Poster based around the Oxy’Pharm Nocospray 2 device and Nocolyse One Shot disinfectant. The world-leading institution achieved repeated Log 6 reductions against Geobacillus in their Containment Level 3 laboratories, which handle human pathogens that can cause serious or life-threatening disease. The HSE mandates a testing regime for operators of CL3 laboratories and insists they use Geobacillus stearothermophilus as their target microorganism, as it is an endospore which is known to be resistant to HPV. Therefore, the testers and the HSE are all but certain the process will eliminate all other pathogens across the Hierarchy of Disinfection. | As Column B. | log5 | Greater than the requirements of the Standards i.e. greater than 4Log10 for spores and greater than 5Log10 for all other organism groups. Phileas has also been independently validated by the UK HSE Laboratory, Buxton, UK. Where it achieved greater than 6log10 reductions across all organism groups. (Report available) | Greater than the requirements of the Standards i.e. greater than 4Log10 for spores and greater than 5Log10 for all other organism groups. Phileas has also been independently validated by the UK HSE Laboratory, Buxton, UK. Where it achieved greater than 6log10 reductions across all organism groups. (Report available) | Greater than the requirements of the Standards i.e. greater than 4Log10 for spores and greater than 5Log10 for all other organism groups. Phileas has also been independently validated by the UK HSE Laboratory, Buxton, UK. Where it achieved greater than 6log10 reductions across all organism groups. (Report available) | Greater than the requirements of the Standards i.e. greater than 4Log10 for spores and greater than 5Log10 for all other organism groups. Phileas has also been independently validated by the UK HSE Laboratory, Buxton, UK. Where it achieved greater than 6log10 reductions across all organism groups. (Report available) | Greater than the requirements of the Standards i.e. greater than 4Log10 for spores and greater than 5Log10 for all other organism groups. Phileas has also been independently validated by the UK HSE Laboratory, Buxton, UK. Where it achieved greater than 6log10 reductions across all organism groups. (Report available) | Effective on Bacteria, Fungi, Viruses, Yeast and Spores Log 5 upto Log 6 reduction
Does the room need evacuating befire the machine is used | Yes | Yes | Yes | Yes | Yes | YES | YES | The ProXcide system is designed to be used in an unoccupied space. | Yes | As Column B. | no | YES. Under UK BPR all automated airborne disinfection must be undertaken in an unattended manner | YES. Under UK BPR all automated airborne disinfection must be undertaken in an unattended manner | YES. Under UK BPR all automated airborne disinfection must be undertaken in an unattended manner | YES. Under UK BPR all automated airborne disinfection must be undertaken in an unattended manner | YES. Under UK BPR all automated airborne disinfection must be undertaken in an unattended manner | Yes
Please advise your injection amount as set out in your testing certification | 8ml of Amity HP 75 per meter3 | from 3 to 8ml/m3 depending on protocol selected (16-18ml/min) | 10g/m3 | 10g/m3 | The equipment is able to deliver fully compliant cycles at 3ml/m3, 6ml/m3 and 8ml/m3 | 0,5 ML/M3 | 0,5 ML/M3 | The ProXcide injects sufficient hydrogen peroxide vapour to ensure the correct parameters for efficacy are achieved on each cycle. This avoids injecting either more than is required, or less than would achieve efficacy. Therefore, there is no pre-set amount for injection. | Between 3 and 5 ml per cubic metre - depending on pathogen being tested against. | As Column B. | N/A | 12ml/cubic metre | 12ml/cubic metre | 12ml/cubic metre | 12ml/cubic metre | 12ml/cubic metre | Flow of 1ml/m3
Please advise your contact time as set out in your testing certification | 60 minutes | From 15 minutes to 120 minutes depending on protocol used | Airborne disinfection contact (ADC) time = 81 mins 30 secs (please note ADC time covers the time period from start of injection to the start of aeration) | Airborne disinfection contact (ADC) time = 81 mins 30 secs (please note ADC time covers the time period from start of injection to the start of aeration) | The equipment is able to deliver fully compliant cycles at 15 minutes, 60 minutes and 120 minutes contact time | 30 MIN | 30 MIN | Automated process - sealf calibrates | Between 1 and 2 hours depending on pathogen being tested against. | As Column B. | 1-5 minutes | 20 minutes | 20 minutes | 20 minutes | 20 minutes | 20 minutes | 2h according to test report norm EN17272
Please advise dwell time as set out in your testing certification | 60 minutes | From 15 minutes to 120 minutes depending on protocol used | 35 mins | 35 mins | The term dwell time is typically used as contact time | ? | ? | Automated process - self calibrates | Total contact time covers dwell time. | As Column B. | N/A | 20 minutes | 20 minutes | 20 minutes | 20 minutes | 20 minutes | 2h according to test report norm EN17272
Please advise aeration time as set out in your testing certification |  | Depends on the room size and protocol, and if adsorber (Cubair-OX) is used | 2 hours to <1.0ppm | 2 hours to <1.0ppm | Aeration time depends on the room volume and will always deliver a compliant cycle. For a 30m3 room, injected at 6ml/m3 with 15 minutes contact time, the aeration time is 44 minutes (down to <1ppm H2O2). | NONE | NONE | Automated process - self calibrates | Aeration does not form a part of the test certification. Typically aeration is between 30 mins and 2 hours depending on whether a Prentative or Curative treatment has taken place. | As Column B. | N/A | Up to 15 minutes depending on room size and extractor used | Up to 15 minutes depending on room size and extractor used | Up to 15 minutes depending on room size and extractor used | Up to 30 minutes depending on room size and extractor used | Up to 30 minutes depending on room size and extractor used | N/A
Please advise the room size as set out in your testing certification | Testing room size was 90m3. Max room size that can be disinfected is 300m3 | Up to 150m3 | 74.4m3 | 74.4m3 | Certification implies a specific testing cycle in an 87m³ room. The equipment is optimised for rooms up to 100m³ | / | / | Automated process - self calibrates | Room sizes vary according to specifics set out with each norm. These are listed on each validation paper and can be provided upon request. | As Column B. | UK cfu testing was 118cbm | 50 cubic metres | 50 cubic metres | 50 cubic metres | 50 cubic metres | 50 cubic metres | 32m3 according to test report norm EN17272
Can your machine be used with people in the room | No | No | No | No | No | NO | NO | No, it is not safe for people to be in the room whilst the system is running | No. | As Column B. | yes | As stated above, UK BPR requires that automated airborne surface decontamination is an unattended process. The nature of the disinfectant when sprayed as Dry Mist or Vapour creates a respiratory, sight and olfactory risk. All disinfectant are toxic and active substances are an enhanced risk when made airborne as a Dry Mist or Vapour. | As stated above, UK BPR requires that automated airborne surface decontamination is an unattended process. The nature of the disinfectant when sprayed as Dry Mist or Vapour creates a respiratory, sight and olfactory risk. All disinfectant are toxic and active substances are an enhanced risk when made airborne as a Dry Mist or Vapour. | As stated above, UK BPR requires that automated airborne surface decontamination is an unattended process. The nature of the disinfectant when sprayed as Dry Mist or Vapour creates a respiratory, sight and olfactory risk. All disinfectant are toxic and active substances are an enhanced risk when made airborne as a Dry Mist or Vapour. | As stated above, UK BPR requires that automated airborne surface decontamination is an unattended process. The nature of the disinfectant when sprayed as Dry Mist or Vapour creates a respiratory, sight and olfactory risk. All disinfectant are toxic and active substances are an enhanced risk when made airborne as a Dry Mist or Vapour. | As stated above, UK BPR requires that automated airborne surface decontamination is an unattended process. The nature of the disinfectant when sprayed as Dry Mist or Vapour creates a respiratory, sight and olfactory risk. All disinfectant are toxic and active substances are an enhanced risk when made airborne as a Dry Mist or Vapour. | No

## W4 - Product Info
all
Product Information |  |  |  |  | Some- please specify below
Tender Name | Environmental Decontamination
Lot Number | 1
Lot Name | Hydrogen Peroxide Vapour Disinfection Systems/Fogging Machines
If there are different responses required for different product categories please specify in additional columns
| Amity | BES Healthcare | Bioquell | Bioquell | Gama | Infinite | Infinite | Inivos | Innotec | Innotec | Sanital Med | UTB | UTB | UTB | UTB | UTB | Wassenburg Medical
Please provide model number/ model name for each system you are submitting onto the framework. | Q82 - Airdecon 200 | Rhea Compact | Bioquell BQ-50 HC Suite (TD091-0000-230V) | Bioquell ProteQ (Sensor Version - TD092-0000-230V) | Scarlet-HPV | VIRO CLEAN EASY - product code: VR100-CL1 | VIRO2 CLEAN SYSTEM - product code: VR200-CL2 | ProXcide™ | 2000.001- Nocospray | 3100.100 - Nocomax | RADS 1000 | Phileas 15 - (P15) | Phileas 25 - (P25) | Phileas 75 - (P75) | Phileas 250 - (P250) | Phileas 285 - (P285) | Wassenberg Glosair
Product Info Questions | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response
Please confirm the type of chemical used in your machine | Hydrogen Peroxide that is 7.5% concentrated (Amity HP 75) | X-CID | 35% Hydrogen Peroxide w/w, CAS: 7722-84-1 | 35% Hydrogen Peroxide w/w, CAS: 7722-84-1 | Peracetic acid boosted hydrogen peroxide | Hydrogen peroxide + ethylic alcool | Hydrogen peroxide + ethylic alcool | Inivos ProXcin™. A bespoke blend of Hydrogen Peroxide at 7.5% concentration. | Ready to use air and surface disinfectant containing either a 6% (Nocolyse Neutral), 12% (Nocolyse One Shot) or 7.9% (Nocolyse Food) Hydrogen Peroxide content. | As Column B | Super oxidised Water | 7.4% concentration hydrogen peroxide | 7.4% concentration hydrogen peroxide | 7.4% concentration hydrogen peroxide | 7.4% concentration hydrogen peroxide | 7.4% concentration hydrogen peroxide | Hydogen Peroxide
Has your machine been tested and certified with any other chemicals | No | n/a | No | No | No | no | no | No, the unique injection system in the ProXcide™ system is specifically designed and tested for use with ProXcin™ | No | As Column B | No | 7.4% hydrogen peroxide/0.14% Peracetic Acid Mix | 7.4% hydrogen peroxide/0.14% Peracetic Acid Mix | 7.4% hydrogen peroxide/0.14% Peracetic Acid Mix | 7.4% hydrogen peroxide/0.14% Peracetic Acid Mix | 7.4% hydrogen peroxide/0.14% Peracetic Acid Mix | No
Please advise the safe exposure limit for these chemicals | 8 hour TWA is 1ppm and 15 minute STEL is 2ppm | 15 Mins at 2ppm and 8hrs at 1ppm | Hydrogen Peroxide CAS: 7722-84-1 UK.EH40 (2011) – Workplace Exposure Limits [WEL]: Long Term Exposure Limit (8hr), Time Weighted Average of 1 ppm, Short Term Exposure Limit 2 ppm | Hydrogen Peroxide CAS: 7722-84-1 UK.EH40 (2011) – Workplace Exposure Limits [WEL]: Long Term Exposure Limit (8hr), Time Weighted Average of 1 ppm, Short Term Exposure Limit 2 ppm | Working level exposure for Hydrogen Peroxide is <1ppm | 2 ppm | 2 ppm | Long-term exposure limit (8-hr TWA reference period): 1ppm Short-term exposure limit (15-minute reference period): 2ppm | In line with the HSE EH40 workplace exposure limits. 2PPM Short term exposure, 1PPM long term exposure. | As Column B | N/A, there is no WEL limit | For both the LTEL is 1ppm for hydrogen peroxide | For both the LTEL is 1ppm for hydrogen peroxide | For both the LTEL is 1ppm for hydrogen peroxide | For both the LTEL is 1ppm for hydrogen peroxide | For both the LTEL is 1ppm for hydrogen peroxide | Limit value (8h) = 1 ppm Limit value (short term) = 2 ppm
Will you be able to offer a remote diagnostic and resolution service free of charge via modem? | Yes | No | No | No | No | after 30 minutes of contact time from the end of the diffusion phase | after 30 minutes of contact time from the end of the diffusion phase | We offer a remote diagnostic and resolution service as part of our usual service level agreeements. | No. We have not found this to be a necessary requirement due to reliability of equipment. | As Column B | No | No | No | No | No | No | No
Please confirm your shutdown mechanism should a person enter the room | The motion sensors will detect movement and stop the cycle immediately. The error message will be displayed on screen that motion was detached. The warning light will also flash. This will also be logged on the traceability report. | Remote Control | Emergency aeration can be activated at any time during the cycle from both inside of the room using the red emergency stop button on the vaporiser unit, or by using the red button on the remote control lectern HMI. Once activated, the system will immediately stop injecting hydrogen peroxide and will start it's emergency aeration protocol, during which the aeration units will switch on to actively break down the hydrogen peroxide vapour in the room. The control lectern will then display an adjusted cycle countdown to advise the length of time required to wait until the room can be checked as safe. The concentration can be checked at any time using the X-am 5100 low-level sensor provided with every BQ-50 to ensure operators, staff and patients are kept safe. | There is a 'hard' Emergency process stop button on the ProteQ base unit (located inside the room), as well as a 'soft' Abort button on the HMI touch screen (outside of the room). If either button is pressed during a cycle which has progressed into the gassing phase, the system will immediately stop injecting hydrogen peroxide vapour and will enter emergency aeration, during which the integrated aeration unit (and any additonal aeration units) will activate to actively break down the hydrogen peroxide vapour in the room. Emergency aeration will continue until the 'Emergency Stop' alarm is acknowledged by an operator. The alarm can only be acknowledged when the Emergency process stop button has been reset/released. The concentration can be checked at any time using a low-level sensor to ensure operators, staff and patients are kept safe. | The equipment is operated from the exterior of the room via a remote control pane which includes a mechanical shutdown button. |  |  | The system will sound an audible alarm if it senses that a person has entered the room during the decontamination process. The system will also automatically enter the deactivation phase to reduce the exposure as quickly as possible. | Firstly, as a part of our implementation training, we ensure each diffusion is set up to ensure maximum security around any access points to the room. Doors are taped and we recommend the use of signs warning of a process taking place. Should a person enter the room - the device can be stopped immediately by using the remote control. | As Column B | N/A, people can be present in the room during use | Phileas can be turned off remotely if it is noticed someone enters the room to limit the concentration of the exposure. However, once any chemical based disinfectant system has started spraying its disinfectant, turning the unit off will not remove the airborne disinfectant, should somebody inadvertently enter the room. Residual disinfectant is also in the room once the spray cycle is over and the dwell time is taking effect. | Phileas can be turned off remotely if it is noticed someone enters the room to limit the concentration of the exposure. However, once any chemical based disinfectant system has started spraying its disinfectant, turning the unit off will not remove the airborne disinfectant, should somebody inadvertently enter the room. Residual disinfectant is also in the room once the spray cycle is over and the dwell time is taking effect. | Phileas can be turned off remotely if it is noticed someone enters the room to limit the concentration of the exposure. However, once any chemical based disinfectant system has started spraying its disinfectant, turning the unit off will not remove the airborne disinfectant, should somebody inadvertently enter the room. Residual disinfectant is also in the room once the spray cycle is over and the dwell time is taking effect. | Phileas can be turned off remotely if it is noticed someone enters the room to limit the concentration of the exposure. However, once any chemical based disinfectant system has started spraying its disinfectant, turning the unit off will not remove the airborne disinfectant, should somebody inadvertently enter the room. Residual disinfectant is also in the room once the spray cycle is over and the dwell time is taking effect. | Phileas can be turned off remotely if it is noticed someone enters the room to limit the concentration of the exposure. However, once any chemical based disinfectant system has started spraying its disinfectant, turning the unit off will not remove the airborne disinfectant, should somebody inadvertently enter the room. Residual disinfectant is also in the room once the spray cycle is over and the dwell time is taking effect. | Via Remote Control
Do you offer trials of products to customers | Yes | Yes | Demonstrations and trials can be arranged. | Demonstrations and trials can be arranged. | Yes, pending availability | Onsite presentation of products offered | Onsite presentation of products offered | Yes, certainly. | Yes, we discuss this in line with the customers requirements. | As Column B | Yes | YES | YES | YES | YES | YES | Yes

## W5 - Delivery & Install
all
Delivery & Installation Information |  |  |  |  | Some- please specify below
Tender Name | Environmental Decontamination
Lot Number | 1
Lot Name | Hydrogen Peroxide Vapour Disinfection Systems/Fogging Machines
If there are different responses required for different product categories please specify in additional columns
| Amity | BES Healthcare | Bioquell | Bioquell | Gama | Infinite | Infinite | Inivos | Innotec | Innotec | Sanital Med | UTB | UTB | UTB | UTB | UTB | Wassenburg Medical
Please provide model number/ model name for each system you are submitting onto the framework. | Q82 - Airdecon 200 | Rhea Compact | Bioquell BQ-50 HC Suite (TD091-0000-230V) | Bioquell ProteQ (Sensor Version - TD092-0000-230V) | Scarlet-HPV | VIRO CLEAN EASY - product code: VR100-CL1 | VIRO2 CLEAN SYSTEM - product code: VR200-CL2 | ProXcide™ | 2000.001- Nocospray | 3100.100 - Nocomax | RADS 1000 | Phileas 15 - (P15) | Phileas 25 - (P25) | Phileas 75 - (P75) | Phileas 250 - (P250) | Phileas 285 - (P285) | Wassenburg Glosair
Delivery & Installation | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response
Please confirm what your maximum lead time will be to supply the machine that you intend to supply through this agreement? | Within 5 to 7 working days, ex stock | 2 Weeks | 3 weeks | 4 weeks | GAMA typically responds to orders from stock (as is the case currently). Typically, equipment is delivered withing 10 working days from purchase order date. Stock levels are advised when a Purchase Order is issued. | 15 days | 15 days | We normally hold stock available for immediate despatch, as well as sufficient stock to build multiple systems in 2-4 weeks. For very large orders, lead time may be up to 12 weeks. | 3-5 working days | 10-14 working days | 2 Working Days from date of order | 14-21 days | 14-21 days | 14-21 days | 14-21 days | 14-21 days | 4 Weeks
Can you offer a service to remove existing equipment? | Yes | Yes | Only if they are Ecolab manufactured units | Only if they are Ecolab manufactured units | Every capital equipment supplier is liable for the removal of its obsolete equipment for safe, sustainable, and environmentally friendly disposal. Should a customer have the requirement for GAMA to remove third party equipment from site this can be provided by GAMA, provided the relevant information related to the equipment is supplied (e.g. safety data sheets, etc). | NO | NO | Yes, certainly. | This is not something we have ever been asked to do with any previous implementations. Redundant or decommissioned equiopment is usually manged by each individual trust in accordance with their own environmental processes on waste management. | As Colomn B | NO | It is unusal for suppliers to allow comeptitors to remove products from a previous client site. Some manufacturers prohibit it. However, if required and subject to the necessary confoimation in requirements for disposal of a competitors equipment we will provide this service. |  |  |  |  | Yes
Is this service at an extra cost | Yes | Yes | No | No | Should third party equipment removal be needed in relation to the sale of GAMA capital equipment, this service carries no extra cost. | NA | NA | No, providing the customer is purchasing our equipment to replace the old systems. | N/A | As Colomn B | N/A | Yes, depending on the equipment. | Yes, depending on the equipment. | Yes, depending on the equipment. | Yes, depending on the equipment. | Yes, depending on the equipment. | No, based on the end user providing a decontamination certificate.
Please provide details of any periods where you are unable to receive or fulfil orders such as factory shutdowns, national bank holidays, holiday periods etc. | Between the Christmas and New Year shut down and National Bank Holidays. | Bank holidays, Warehouse Shutdown between Christmas/NY and Annual stock take (first week in april) | The office and factory are shut on all public holidays and the weekends. | The office and factory are shut on all public holidays and the weekends. | GAMA processes and fulfil orders during normal UK working periods (Monday to Friday, 9am to 5am). | 24/7 365 service. Ordered processed everyday of the year. | 24/7 365 service. Ordered processed everyday of the year. | We maintain staffing for the entire year. | National Bank Holidays | As Colomn B | We are unable to fulfil orders on public holidays and bank holidays | National Bank Holidays only | National Bank Holidays only | National Bank Holidays only | National Bank Holidays only | National Bank Holidays only | N/A

## W6 - Training & Safety
Technical Training Information
Tender Name | Environmental Decontamination
Lot Number | 1
Lot Name | Hydrogen Peroxide Vapour Disinfection Systems/Fogging Machines
If there are different responses required for different product categories please specify in additional columns
| Amity | BES Healthcare | Bioquell | Bioquell | Gama | Infinite | Infinite | Inivos | Innotec | Innotec | Sanital Med | UTB | UTB | UTB | UTB | UTB | Wassenburg Medical
Please provide model number/ model name for each system you are submitting onto the framework. | Q82 - Airdecon 200 | Rhea Compact | Bioquell BQ-50 HC Suite (TD091-0000-230V) | Bioquell ProteQ (Sensor Version - TD092-0000-230V) | Scarlet-HPV | VIRO CLEAN EASY - product code: VR100-CL1 | VIRO2 CLEAN SYSTEM - product code: VR200-CL2 | ProXcide™ | 2000.001- Nocospray/3100.100 - Nocomax | 2000.001- Nocospray/3100.100 - Nocomax | RADS 1000 | Phileas 15 - (P15) | Phileas 25 - (P25) | Phileas 75 - (P75) | Phileas 250 - (P250) | Phileas 285 - (P285) | Wassenburg Glosair
Training | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response
How many hours of free of charge on-site, face-to-face clinical end user training, following validation of the equipment, be delivered to accommodate the needs of the customer free of charge ? This must include all training on the operation of the machine. | 16 hours | 1 day | Half day for up to 5 people per unit purchased | Half day for up to 5 people per unit purchased | An implementation process is included with the sale of every GAMA capital equipment item. The implementation period includes face to face training, provided by a dedicated Product Specialist team. Subsequent training is available as required by the customer. All training (initial and subsequent) is provided face to face and is always free of charge. | 3 training sessions of 8 hours each (total 24 training hours) ooffered FOC | 3 training sessions of 8 hours each (total 24 training hours) ooffered FOC | Unlimited training will be provided for all operators when machines are purchased new for the first time. This includes up to 8 hours on site as required. Additional operator training thereafter will be covered under our usual service level agreements | Within reason, we generally liaise with each end user and assess and discuss their individual training requirements. Our aim is to ensure as many personnel as the end user wishes to be trained, are trained free of charge. Depending on numbers, this is routinely covered with either a half or full days training. Generally speaking, the larger the amount of equipment purchased, the greater the need for larger trainee numbers, so we can accommodate this accordingly. | Within reason, we generally liaise with each end user and assess and discuss their individual training requirements. Our aim is to ensure as many personnel as the end user wishes to be trained, are trained free of charge. Depending on numbers, this is routinely covered with either a half or full days training. Generally speaking, the larger the amount of equipment purchased, the greater the need for larger trainee numbers, so we can accommodate this accordingly. | For 1 to 5 machines a maximum of 2 Hours for 5 to 10 machines 4 Hours and for 10+ machines 6 Hours. However, online video tutorials are also freely available 24/7 for regular review as needed. | All initial training for clinical end-users is free of charge. This is availabl to A. Operators - max 10 per session and B. 'Train the Trainers' - maximum 6 per session. We offer a total of 32 hours of free training (4 working days). Phileas units have the same software and deployment requirements. Therefore, training covers deployement of all units. Any additional training for new Opertators and Trainers, beyond this initial period is chargeable. Up to 10 Operators - £1875 +VAT. Train the Trainers uo to 6 people £1500 +VAT |  |  |  |  | 2 x days i.e. from 09:00 a.m. > 17:00 p.m. Additional days can be supplied at no additional cost
Please confirm whether your user manual includes the following: - basic machine operation - basic trouble shooting - contact details incase of problems | -Yes -Yes -Yes | Yes | The user manual includes all aspects listed | The user manual includes all aspects listed | All of these aspects are included in the user manual. | Yes | Yes | All points confirmed. | Yes, we have a very comprehensive training manual and operational guide which highlight these areas | Yes, we have a very comprehensive training manual and operational guide which highlight these areas | Yes / Yes / Yes | YES | YES | YES | YES | YES | Yes
Can the time allocation of any free of charge, on-site, face-to-face clinical end user training, following the validation of the equipment, be delivered to accommodate the needs of the customer ? i.e it can be delivered weekends, for part-days and outside the hours of Mon to Fri, 9 a.m. -5 p.m. (GMT/BST), on a free of charge basis. If you do not offer free of charge face-to-face training please select `No` to this question, your response must refer to all products. | Yes | Free of Charge Face to face training Mon-Fri 9am-5pm only | No | No | Training is typically delivered face to face during working days. Night shift training is on occasion delivered pending customer request and availability. All training is free of charge. | Yes | Yes | Yes | Generally speaking, we deliver the majority of training requirements Monday to Friday between 9am - 5pm. We very much encourage a 'train the trainer' approach where nominated senior personnel are trained to a point where they can pass on this knowledge to others. This not only assists in ensuring those working outside of these hours receive their training, but also assists when newly recruited personnel join the team and require training too. We liaise with each individual end user and assess and discuss their training needs and formulate a training plan ahead of their purchase. | Generally speaking, we deliver the majority of training requirements Monday to Friday between 9am - 5pm. We very much encourage a 'train the trainer' approach where nominated senior personnel are trained to a point where they can pass on this knowledge to others. This not only assists in ensuring those working outside of these hours receive their training, but also assists when newly recruited personnel join the team and require training too. We liaise with each individual end user and assess and discuss their training needs and formulate a training plan ahead of their purchase. | Outside of working hours to cover shifts from Mon to Fri. Weekends are available on a charged basis. | Part days are acceptable, but out-of-hours training is not available. For initial training we will accommodate weekends if required. |  |  |  |  | No
Is new employee technical training offered? | Yes | Yes | Training can be purchased and arranged at any time following the installation of the unit | Training can be purchased and arranged at any time following the installation of the unit | The training provided is for the safe and effective operation of the device. All technical support is provided by GAMA’s technical team and not by the end user. | Yes | Yes | Yes | Yes - via a 'train the trainer' approach or ad hoc training agreed between ourselves and the end user. | Yes - via a 'train the trainer' approach or ad hoc training agreed between ourselves and the end user. | Providing it falls within the free hours provided, otherwise chargeable however, online video tutorials are freely available to help introduce new employees to the system. | YES | YES | YES | YES | YES | Yes
Is off site pre delivery/post delivery technical training available? | Yes | Yes | Training can be carried out at our site in Andover | Training can be carried out at our site in Andover | No end user technical training is currently available. All technical maintenance of the equipment is provided by GAMA. | Yes | Yes | Yes | Yes - we endevour to deliver training to the needs of the end user wherever possible. | Yes - we endevour to deliver training to the needs of the end user wherever possible. | No | Not at the moment, it is under consideration | Not at the moment, it is under consideration | Not at the moment, it is under consideration | Not at the moment, it is under consideration | Not at the moment, it is under consideration | Yes
Is there a cost involved in 'off-site' training? | No | Yes | Yes | Yes | No. Majority of our training is provided on site as per customer’s requirements but we have a training room in our North & South bases in the UK. Training is provided free of charge | No | No | Yes | This would need to be assessed on an individual basis depending on where the off-site training takes place and what costs may be involved. We liaise with each individual end user and assess and discuss their training needs and formulate a training plan ahead of their purchase. | This would need to be assessed on an individual basis depending on where the off-site training takes place and what costs may be involved. We liaise with each individual end user and assess and discuss their training needs and formulate a training plan ahead of their purchase. | N/A | Not applicable at this time | Not applicable at this time | Not applicable at this time | Not applicable at this time | Not applicable at this time | Yes
If yes detail the costs inc VAT | N/A | TBC | 1317.6 | 1317.6 | N/A |  |  | To be confirmed subject to requirements | This is unlikely to be a standard fixed cost and so will need to be assessed on an individual basis depending on where the training takes place and what the resepctive costs may be. | This is unlikely to be a standard fixed cost and so will need to be assessed on an individual basis depending on where the training takes place and what the resepctive costs may be. | N/A | Not applicable at this time | Not applicable at this time | Not applicable at this time | Not applicable at this time | Not applicable at this time | £750 per day upto 6 users
Do you provide technical manuals with the equipment? | Yes | Yes | All information needed is included in the Operation and Maintenance Manual | All information needed is included in the Operation and Maintenance Manual | No. All technical support is provided by GAMA. | Yes | Yes | No | Yes | Yes | Yes | YES | YES | YES | YES | YES | Yes
- When providing the technical manuals please advise what format these are presented to the trust in - Hard copy/ online/ CD | Hard copy, CD, USB, email, on the machine remote | Hard Copy | On line | On line | Hard copy/ online/ CD | Hard Copy, Online, E learning | Hard Copy, Online, E learning | N/A | Hard Copy, pdf and on line video. | Hard Copy, pdf and on line video. | Hard Copy & Online | Hard copy and CD | Hard copy and CD | Hard copy and CD | Hard copy and CD | Hard copy and CD | Hard Copy/On - Line
Please advise if refresher training is available Free of Charge if there is any update to standards/ equipment/ software? | Yes | Annual Training given (if they have a contract) | Yes | Yes | N/A | Yes | Yes | Yes, under the usual service level agreement | From time to time there may be amendments to documentation such as MSDS which is then sent on to the end user to replace the previous edition. End users are also informed of any new products or modifications of existing equipment. Such amendments have never required retraining, as these are deemed to be administrative. | From time to time there may be amendments to documentation such as MSDS which is then sent on to the end user to replace the previous edition. End users are also informed of any new products or modifications of existing equipment. Such amendments have never required retraining, as these are deemed to be administrative. | Yes | Refresher training is available, free of charge, for updates etc. to the 'Train the Trainer' groups
Is this training CPD accredited? | No | No | No | No | N/A | NO | NO | Yes | The training of our equipment does not require an accreditation of this level. We have implemented (including training) of more than 350 of these machines within the NHS - we pride ourselves that each and every implementation has been carried out to the total satisfaction of the respective end user both in terms of training activity and after training supporting materials. | The training of our equipment does not require an accreditation of this level. We have implemented (including training) of more than 350 of these machines within the NHS - we pride ourselves that each and every implementation has been carried out to the total satisfaction of the respective end user both in terms of training activity and after training supporting materials. | No | No | No | No | No | No | No.
Safety | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response |  | Supplier Response | Supplier Response
Please advise how you would prepare the room before using the product | remove all waste annd linan, clean all surfaces with a disinfectant and position items within the room to maximise hydrogen peroxide exposures | Internal ventilation and room sealed | No-touch decontamination processes, including the Bioquell vapour decontamination process, should not replace or remove the need for manual cleaning. Therefore, our recommendaiton is for the target room to be pre-cleaned to remove gross contamination commencing the Bioquell decontamination process. Absorbant materials should be removed from the area (e.g. bed linens, upholstered items, paper towels etc.), and as many surfaces exposed as possible by moving furniture and opening cupboard doors, drawers etc. Upon exit, the door to the room should be sealed with tape, and a no entry warning sign placed on the door. | No-touch decontamination processes, including the Bioquell vapour decontamination process, should not replace or remove the need for manual cleaning. Therefore, our recommendaiton is for the target room to be pre-cleaned to remove gross contamination commencing the Bioquell decontamination process. Absorbant materials should be removed from the area (e.g. bed linens, upholstered items, paper towels etc.), and as many surfaces exposed as possible by moving furniture and opening cupboard doors, drawers etc. Upon exit, the door to the room should be sealed with tape, and a no entry warning sign placed on the door. | Room preparation from an efficacy standpoint: To optimise the efficacy of the process, surfaces must be exposed to allow access to the HPV aerosol, occluded surfaces will not be disinfected. This includes practices such as moving the mattress on to its side to expose the bed frame, open cupboard doors and drawers etc. |  |  | Detailed training will be provided, but main points are as follows: - Ensure the room is unoccupied - Cover ventilation grilles and smoke detectors using Inivos VentKits and CapKits (no ladders required) - Place the system in the room and connect the power - Remove absorbent items (curtains and mattresses. - Seal the room using Inivos DVI sealant tape to prevent vapour egress - Room is ready to be decontaminated | It is essential that a robust manual cleaning process in line with the end user's protocols are carried out prior to the use of our equipment. The room must be able to be completely sealed by closing windows and doors. It is recommended that smoke alarms and air vents are covered (Cover Kit available via Supply Chain). Where appropriate, open cupboards and drawers. If a bed mattress is in the room, lift this up and lie on bed frame to give maximum exposure to diffused gas.Ensure noone is in the room at point of commencing the disinfection procedure. | It is essential that a robust manual cleaning process in line with the end user's protocols are carried out prior to the use of our equipment. The room must be able to be completely sealed by closing windows and doors. It is recommended that smoke alarms and air vents are covered (Cover Kit available via Supply Chain). Where appropriate, open cupboards and drawers. If a bed mattress is in the room, lift this up and lie on bed frame to give maximum exposure to diffused gas.Ensure noone is in the room at point of commencing the disinfection procedure. | Regular Cleaning protocols are recommended prior to use | See 30 below | See 30 below | See 30 below | See 30 below | See 30 below | Isolate the room
Please advise all suitable PPE you would recommend for using this equipment | Coverall Type 5 and 6 (Meets EN 14605 or eqivalent) Tape for sealing the room (3M PC- 3434) Goggles - None vented clear (Meets EN 166 I.F.N or equivalent) Mask with cartridge (3M PC 4251 & 13014) Gloves Apron | Standard PPE | When handling hydrogen peroxide you must always were gloves, eye protection and long sleeves. PPE should not be required during system cycle operation due to the use of catalytic aeration units to remove the hydrogen peroxide down to below the safe exposure limit, as confirmed with the low-level sensor provided. | When handling hydrogen peroxide you must always were gloves, eye protection and long sleeves. PPE should not be required during system cycle operation due to the use of catalytic aeration units to remove the hydrogen peroxide down to below the safe exposure limit, as confirmed with the low-level sensor provided. | At no point the operator should be in the room when the machine is in operation or when the ambient levels of hydrogen peroxide (H2O2) are beyond what is stipulated by the Health and Safety Executive (HSE). The Working Exposure Limit (WEL) for H2O2 is <1ppm, only reached when the cycle is concluded. An external display will inform the operator when the cycle is complete. For these reasons there is required for PPE for routine daily use of clinell scarlet. It is advised for goggles and gloves to be used when replacing the liquid biocide bottle in the machine, although the risk is minimal as no pouring etc is required. | Basic PPE as per HSE | Basic PPE as per HSE | No PPE is stipulated, but it may be recommended to use gloves and safety goggles when handling the chemical bottles. | Gloves, Sealed eye protection and Filtering face piece (FFP) face mask. FFP classification as: FFP1 (post contact time entry - see operational guide) FFP3 (during or immediately after diffusion). | Gloves, Sealed eye protection and Filtering face piece (FFP) face mask. FFP classification as: FFP1 (post contact time entry - see operational guide) FFP3 (during or immediately after diffusion). | Not Required | Portable gas monitor with hydrogen peroxide sensor for a 0-100ppm range. Our recommendation is ATi Portasens III. 3M full/half face respirators suitable for 3m 6003/6006 filter cartridges for room re-entry, these are validated for hydrogen peroxide, up to 120 hours use at 90ppm. If Half-Face mask selected then suitable, sealed goggles must be worn at the same time for eye protection. Disposable gloves and aprons for handling disinfectant containers. 2 inch masking tape for door sealing. Barrier tape for all public access doors. Highly visible No-Entry signage for all entry doors. plastic sheeting for grille covering. Extendable ladders/steps for covering high grilles. extendable poles and covers for ceiling grilles. Extension leads for optimal placing of Phileas and extractor units. |  |  |  |  | Gloves,mask FFP3 and goggles should be worn when handling the liquid Hydrogen Peroxide H2O2
Please advise any room preparations that are required i.e.sealing to ensure no one will enter while in use | use tape across any gaps to prevent the gas from escaping. Place tape across the door to prevent anyone from entering. Use the signage that is provided with machine and place it outside the room. Cover the fire alarm and seal all air vents. | Internal ventilation and room sealed & warning sign on the door | The target room must be sealed to prevent the escape of vapour into neighbouring areas, this includes the sealing of heating, ventilation and air conditioning (HVAC) vents within the target room. In order to make room set-up as fast as possible, two Bioquell VS-1 products are included with each BQ-50 suite, which can be used for the rapid closure of most ceiling vents. For areas where the Bioquell VS-1 can't be used, alternative sealing arrangements need to be made (this is covered in the Bioquell training documentation and operator training). | The target room must be sealed to prevent the escape of vapour into neighbouring areas, this includes the sealing of heating, ventilation and air conditioning (HVAC) vents within the target room. In order to make room set-up as fast as possible, Bioquell VS-1 products can be used for the rapid closure of most ceiling vents. For areas where the Bioquell VS-1 can't be used, alternative sealing arrangements need to be made (this is covered in the Bioquell training documentation and operator training). | Room preparation from a safety standpoint: It is critical to contain the treated room to ensure the HPV does not migrate into neighbouring rooms. This safety requirement is universal across any supplier. Whilst there is no pressure build up as part of the process it is important to ensure that there are no physical pathways for the HPV to escape. This will require for ventilation inlets/outlets to be blocked, ensure false ceilings are complete (e.g. no missing ceiling tiles), close and seal doors with non-porous tape, etc. A portable H2O2 monitor is issued by default with the machine in order to monitor any potential leaks onto corridors etc. | None | None | These points are covered in the room preparation above. | Internal room preparations are indicated above. External preparations include the taping around the edges of the operator's exit door and signage warning not to enter due to decontamination process being carried out. | Internal room preparations are indicated above. External preparations include the taping around the edges of the operator's exit door and signage warning not to enter due to decontamination process being carried out. | No | Rooms must be pre-cleaned to remove all visible dirt. All porous materials need to be removed. All food stuffs need to be removed or safely locked away. If possible, smoke detectors need to be turned off. if they cannot be turned off they must be suitably covered. All cupboards need to be opened. Mattresses and pillows need to be clear of the bed. No plants in the room. Doors or grilles leading to other areas need to be sealed and covered to prevent leaks, and warning signs posted. with Phileas unit and extractor positioned, access door is closed, sealed, taped and No Entry signage posted. MyPhileas Tablet is linked to the unit to start the cycle from outside the sealed room. |  |  |  |  | Sealing of room to ensure no one enters or exits room while in use

## W7 - Warranty
some- please state below
Warranty Information |  |  |  |  | all
Tender Name | Environmental Decontamination
Lot Number | 1
Lot Name | Hydrogen Peroxide Vapour Disinfection Systems/Fogging Machines
| AMITY | BES Healthcare | Bioquell | Bioquell | Gama | Infinite | Infinite | Inivos | Innotec | Innotec | Sanital med | UTB | UTB | UTB | UTB | UTB | Wassenburg Medical
Please provide model number/ model name for each system you are submitting onto the framework. | Q82 - Airdecon 200 | Rhea Compact | Bioquell BQ-50 HC Suite (TD091-0000-230V) | Bioquell ProteQ (Sensor Version - TD092-0000-230V) | Scarlet-HPV | VIRO CLEAN EASY - product code: VR100-CL1 | VIRO2 CLEAN SYSTEM - product code: VR200-CL2 | ProXcide™ | 2000.001- Nocospray/3100.100 - Nocomax | 2000.001- Nocospray/3100.100 - Nocomax | RADS 1000 | Phileas 15 - (P15) | Phileas 25 - (P25) | Phileas 75 - (P75) | Phileas 250 - (P250) | Phileas 285 - (P285) | WASSENBURG GLOSAIR
System Specific Information | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response | Supplier Response
Please confirm if any of your current core range is due to be discontinued within the next 6 years | Range extensions will be happening but no decision has been made to discontinue the existing range | No | No | No | Given the market requirement for innovation in this sector and GAMA’s ability to develop and launch innovative products it is possible that the current equipment on offer may be upgraded within the next 6 years. | NO | NO | No | No discontinuation is planned. | No discontinuation is planned. | NO | No | No | No | No | No | No
From the point of discontinuance of a machine, how long will you continue to provide spare parts for maintenance and repair of the obsolete machine? | We would aim to supply spare parts up to 10 years | 1 year | 5 years | 5 years | Obsolete equipment will be maintained for a minimum period of four years post discontinuation. Alternatively, customers have the option for preferential rates to upgrade to current specification equipment. | 7 Years | 7 Years | 6 Years | Minimum of 5 years. | Minimum of 5 years. | 2 Years | As required by law for electrical products, up to 7 years | As required by law for electrical products, up to 7 years | As required by law for electrical products, up to 7 years | As required by law for electrical products, up to 7 years | As required by law for electrical products, up to 7 years | Should the machine be discontinued, spare parts will be supplied for up to 10 years.
Within the warranty period
When will your Customer Service Helpdesk be available outside of the stated hours specified in the Specification (Mon to Fri, 9 a.m. -5 p.m. (GMT/BST)? If your Customer Service Helpdesk is only avaliable Mon to Fri, 9 a.m. -5 p.m. (GMT/BST), your response will attract a ZERO score as this is the minimum requirement. | Trusts will have access to a 24 hour phone line | Only Mon-Fri 8.30-5 | No | No | Customer service helpdesk is available only 9am-5pm BST. An emergency contact outside of these hours is available – Global Technical Sales Manager | 24/7 Customer Service | 24/7 Customer Service | We operate a full time helpdesk service. 24hrs a day, 7 days a week, 365 days a year. | Our customer service telephone number is available 24 hours a day. | Our customer service telephone number is available 24 hours a day. | Available (Mon to Fri 7am- 6pm GMT/BST) | A 24/7 Helpdesk number will be available that will always be manned. | A 24/7 Helpdesk number will be available that will always be manned. | A 24/7 Helpdesk number will be available that will always be manned. | A 24/7 Helpdesk number will be available that will always be manned. | A 24/7 Helpdesk number will be available that will always be manned. | No
In the event of any equipment failure, from the point you are notified, how long will it take for an engineer to be on site during standard working hours? Standard working hours are defined as Mon to Fri 9 a.m. to 5 p.m. (GMT/BST) excluding bank holidays. | Normally within 2 business days | Upto 5 working days but usually 24-72Hrs | 2 working days (18 standard working hours) | 2 working days (18 standard working hours) | Standard service level agreements consist of a 48 working hour lead time for a site visit. | 6-8 hours | 6-8 hours | 24 hours from receipt of call, 24/7 | We endeavour to ensure an engineer is on site within 72 working hours of a report of failure | We endeavour to ensure an engineer is on site within 72 working hours of a report of failure | 2-days Following day loan machine service available. | The self-diagnosis properties of the Phileas software allow a decision to be made on the phone whether a call-out is required. Call-outs where replacement equipment is agreed with the client will be within 4 hours during a standard working day. Calls outside standard working days will be responded to no later than the next working day. |  |  |  |  | This is dependent on the contract status
Please confirm the number of service visits included in the warranty period (per annum) | 1 per annum scheduled service visit | 0 | No maximum for warranty issues | No maximum for warranty issues | The number of site visits related to (statutory or extended) warranty claims is unlimited. | 1 | 1 | 4 no. proactive visits, unlimited reactive visits, under usual SLA agreement | It is recommended that the machines are serviced and calibrated annually. A standard service charge will be applied. | It is recommended that the machines are serviced and calibrated annually. A standard service charge will be applied. | One | Warranty period is one year. However, pricing includes 3 service visits to be undertaken annually from year 1. | Warranty period is one year. However, pricing includes 3 service visits to be undertaken annually from year 1. | Warranty period is one year. However, pricing includes 3 service visits to be undertaken annually from year 1 | Warranty period is one year. However, pricing includes 3 service visits to be undertaken annually from year 1. | Warranty period is one year. However, pricing includes 3 service visits to be undertaken annually from year 1. | 1
Are you able to provide an engineer on site for a system malfunction call-out on the same day or on the next working day (Mon to Fri 8am to 6pm) during the warranty period? | Yes | Yes | During standard working hours (Mon to Fri 9am to 5pm) we will be on site in 48 hours, where possible this will be the next day | During standard working hours (Mon to Fri 9am to 5pm) we will be on site in 48 hours, where possible this will be the next day | Lead times for site visits are 48 working hours. The majority of requests are met within 24 working hours | Same Day | Same Day | Yes | We endeavour to ensure an engineer is on site as soon as possible following a reported failure - this is irrespective of whether the machine is under warranty or not. Same or next day service will be strived for if required. | We endeavour to ensure an engineer is on site as soon as possible following a reported failure - this is irrespective of whether the machine is under warranty or not. Same or next day service will be strived for if required. | Yes | Yes | Yes | Yes | Yes | Yes | Yes
How many trained engineers cover this equipment in England | 4 | 2 | 10 | 10 | Four field technicians cover England. A further three Product Specialists will address issues related to lack of training and consumables (e.g. sensors, lamp failures) | 4 | 4 | 12 | 6 | 6 | 1 per 100 machines | 4 | 4 | 4 | 4 | 4 | 30
Are all engineers who will be servicing this equipment employed by yourselves? | Yes | Yes | Yes | Yes | Yes. All GAMA equipment is supported directly by GAMA staff with regards to sales, training, implementation, maintenance and repairs. | Y | Y | Yes | No. We out source to our service and maintenance partner - CK Consumables. | No. We out source to our service and maintenance partner - CK Consumables. | Yes | YES | YES | YES | YES | YES | Yes
If NO please state their employer and terms of the agreement to service the equipment | N/A | n/a |  |  | N/A |  |  |  | CK Consumables are our service and maintenance partner who employ the engineers. They have been servcing and calibrating our equipment since 2012. | CK Consumables are our service and maintenance partner who employ the engineers. They have been servcing and calibrating our equipment since 2012. | N/A | NO | NO | NO | NO | NO
Are engineers, who are required to visit customer sites subject to DBS Checks/Criminal Records? If YES please give details | No | Yes | No | No | Yes, all GAMA field staff is DBS checked. | Yes | Yes | Yes - DBS and Criminal records | TBA | TBA | Yes Basic Disclosure | NO | NO | NO | NO | NO | Yes - Disclosure Barring Services - Basic checks
Will you be able to offer a remote diagnostic and resolution service free of charge via modem? | Yes | No | No | No | No, the current equipment does not have the ability to share data remotely. Data is collected in the device locally which can be extracted via USB by the customer (as per the training) and emailed by the customer should there be value to diagnose a given fault. | Yes | Yes | Yes, under the usual service level agreement | No. Our equipment is relatively simplistic in terms of its composition and does not require this functionality. | No. Our equipment is relatively simplistic in terms of its composition and does not require this functionality. | No | No. Phileas provides extensive information on its local reporting. There is nothing that could be suitably resolved remotely | No. Phileas provides extensive information on its local reporting. There is nothing that could be suitably resolved remotely | No. Phileas provides extensive information on its local reporting. There is nothing that could be suitably resolved remotely | No. Phileas provides extensive information on its local reporting. There is nothing that could be suitably resolved remotely | No. Phileas provides extensive information on its local reporting. There is nothing that could be suitably resolved remotely | No
What is the average delivery lead time for spare parts within the last 12 months | 7 working days | 2 Weeks | For critical spares next working day | For critical spares next working day | For devices covered by a statutory or extended warranty agreement the standard SLA quotes a lead time for repair completion of 72 hours. Should this not be possible a replacement (identical) unit is provided until the repair is completed. | 24 hours | 24 hours | 1-2 days | External spare parts are kept in stock and can be dispatched the same or next day depending on time of order. Internal parts can only be fitted by a trained engineer and these parts are routinely stocked in their van. | External spare parts are kept in stock and can be dispatched the same or next day depending on time of order. Internal parts can only be fitted by a trained engineer and these parts are routinely stocked in their van. | 2 working days | Spare parts will be kept as stock. Suitbale spare parts will be available within 24-48 hours | Spare parts will be kept as stock. Suitbale spare parts will be available within 24-48 hours | Spare parts will be kept as stock. Suitbale spare parts will be available within 24-48 hours | Spare parts will be kept as stock. Suitbale spare parts will be available within 24-48 hours | Spare parts will be kept as stock. Suitbale spare parts will be available within 24-48 hours | 4 weeks
Based on the repair completion time in >95% of instances in the last 12 months, please confirm your repair completion time (In Working Days) for minor repairs of machines. | 1 day | 1-2 Days | 1 day | 1 day | Minor repairs (e.g. filter replacement) have been met withing 48 working hours in >95% of cases | 1 working day | 1 working day | 1-2 days | Most minor repairs are completed on the day of engineers visit. | Most minor repairs are completed on the day of engineers visit. | 2 working days | Less than 1 working day. All Phileas units share the same components and software making repairs simple and efficient. | Less than 1 working day. All Phileas units share the same components and software making repairs simple and efficient. | Less than 1 working day. All Phileas units share the same components and software making repairs simple and efficient. | Less than 1 working day. All Phileas units share the same components and software making repairs simple and efficient. | Less than 1 working day. All Phileas units share the same components and software making repairs simple and efficient. | 1-2 days
Based on the repair completion time in >95% of instances in the last 12 months, please confirm your repair completion time (In Working Days) for major repairs of machines. | 2 days | 1-2 Days | 1 day - for major repairs a replacement unit would be sent out during the warranty period while your unit is fixed | There were no major repairs required in the last 12 months | Major repairs have been completed under 72 hours or a replacement machine has been provided in >95% of cases. | 2-3 working days | 2-3 working days | 4-8 days | Instances of major repair are rare. However, where this may be apparent and an internal rebuild may be required, this may take up to 5 working days. | Instances of major repair are rare. However, where this may be apparent and an internal rebuild may be required, this may take up to 5 working days. | 2 working days | Less than 1 working day. All Phileas units share the same components and software making repairs simple and efficient. | Less than 1 working day. All Phileas units share the same components and software making repairs simple and efficient. | Less than 1 working day. All Phileas units share the same components and software making repairs simple and efficient. | Less than 1 working day. All Phileas units share the same components and software making repairs simple and efficient. | Less than 1 working day. All Phileas units share the same components and software making repairs simple and efficient. | 1-2 days
What will the minimum "guaranteed uptime" be across the range of products that you are tendering for under this Lot ? Guaranteed uptime refers to the amount of time the products are fully operational excluding time for daily, weekly and periodic / service tests. | 95% | 95% | There is no guaranteed up time, however the 2021 average time between failures was 62 weeks | There have been no failures in 2021 of the installed based in the UK | Uptime data has not been formally collected. Downtime related with preventative maintenance is minimal, circa one hour per 12 months. | 98% | 98% | 480 hours | Nocotech machines require an annual service and calibration, which means the device will be available for use throughout the rest of the year. The systems are extremely reliable, with few moving parts, and most downtime is usually down to misuse. In such circumstances, we endeavour to effect the repair with the shortest lead time possible. | Nocotech machines require an annual service and calibration, which means the device will be available for use throughout the rest of the year. The systems are extremely reliable, with few moving parts, and most downtime is usually down to misuse. In such circumstances, we endeavour to effect the repair with the shortest lead time possible. | 95% | Minimum guaranteed uptim is greater than 95% for normal deployment and useage, as recommended. | Minimum guaranteed uptim is greater than 95% for normal deployment and useage, as recommended. | Minimum guaranteed uptim is greater than 95% for normal deployment and useage, as recommended. | Minimum guaranteed uptim is greater than 95% for normal deployment and useage, as recommended. | Minimum guaranteed uptim is greater than 95% for normal deployment and useage, as recommended. | We are unable to confirm the below information, and this is premarily that WBUK have just recently become the UK agent for this product. We are currently assessing the installed base and putting in place service contracts to support the installed base, at this time we do not have the "guranteed uptime". We will update this once we have the information.
What remidies are available to the hospital if Guaranteed Up-time is not achieved | if due to machine fault, we offer a swap in, swap out service to help ensure trust up time. | Work with the trust to pinpoint the reasons why the uptime wasn’t achieved and put processes in place to rectify this. | A priority service contract is available after the warranty which includes loan units to ensure that up time is maintained | A critical service contract is available after the warranty period which includes part stock holding on site and a 1 working day response time to breakdowns | A replacement unit can be issued to ensure continuity of service. | machine replacement if due to a fault in the machine. Nor covered for accidental or operational errors | machine replacement if due to a fault in the machine. Nor covered for accidental or operational errors | Immediate attendance on site, swap out equipment provided until repair is made | Loan devices are available on request, but may be recharged to the client if repairs fall outside the standard warranty. | Loan devices are available on request, but may be recharged to the client if repairs fall outside the standard warranty. | Any issues we will swap out the machine | Any unit failing to meet this criteria, will be replaced. | Any unit failing to meet this criteria, will be replaced. | Any unit failing to meet this criteria, will be replaced. | Any unit failing to meet this criteria, will be replaced. | Any unit failing to meet this criteria, will be replaced. | 4-5 hours.
What down time allowance is required for the quarterly & annual service visit? | 1 day | Annually/Quarterly | Half day | Half day | Downtime related with preventative maintenance is minimal, circa one hour per 12 months. |  |  | Replacement systems will be provided and systems will be returned to our production facility for full service, therefore there will be no service downtime. | Most service and calibration processes are carried out between 1 and two hours per machine. | Most service and calibration processes are carried out between 1 and two hours per machine. | 1 Hour | Phileas only requires an Annual Service visit. Down time is 3-5 hours deplending on the unit | Phileas only requires an Annual Service visit. Down time is 3-5 hours deplending on the unit | Phileas only requires an Annual Service visit. Down time is 3-5 hours deplending on the unit | Phileas only requires an Annual Service visit. Down time is 3-5 hours deplending on the unit | Phileas only requires an Annual Service visit. Down time is 3-5 hours deplending on the unit | Yes
Will all equipment and accessories found to be defective within the warranty period, be replaced or repaired free of charge? | Yes | Yes | Yes subject to normal use in accordance with the operation and maintenance manual | Yes subject to normal use in accordance with the operation and maintenance manual | Yes, all faults in equipment covered by a warranty agreement (statutory or extended by the customer) are remediated free of charge | Replaced | Replaced | Yes | Defective parts will be replaced/repaired free of charge providing there is no apparent misuse of the equipment or wilful damage. | Defective parts will be replaced/repaired free of charge providing there is no apparent misuse of the equipment or wilful damage. | Yes | Yes, subject to normal, recommended use. Obvious abuse and non-recommended use will be chargeable. | Yes, subject to normal, recommended use. Obvious abuse and non-recommended use will be chargeable. | Yes, subject to normal, recommended use. Obvious abuse and non-recommended use will be chargeable. | Yes, subject to normal, recommended use. Obvious abuse and non-recommended use will be chargeable. | Yes, subject to normal, recommended use. Obvious abuse and non-recommended use will be chargeable. | Free of charge labour and parts
Please detail what is included in the warranty | AmityInternationalundertake that if, within the warranty period, our equipment or any part thereof, is proved to be defective by reason only of faulty workmanship or materials, we will at our option, repair or replace the same free of charge. | Manufacturing defects | One year warranty from after installation or within 30 days of delivery which ever occurs sooner, including all parts and labour. It is conditional on the basis of normal operation in accordance with the operation and maintenance manual. Filters are not covered by the equpment warranty. | One year warranty from after installation or within 30 days of delivery which ever occurs sooner, including all parts and labour. It is conditional on the basis of normal operation in accordance with the operation and maintenance manual. Filters are not covered by the equpment warranty. | A predefined period during which GAMA guarantees its capital equipment to perform and remain in the originally specified condition, by carrying out technical repairs or equipment exchanges. |  |  | Full details of the warranty provision can be found in our terms and conditions, which are available on request, or can be saved to the online NHSSC catalogue | Any fault with the appliance that is a result of defective parts or workmanship within 2 years of date of purchase will be repaired or replaced free of charge. | Any fault with the appliance that is a result of defective parts or workmanship within 2 years of date of purchase will be repaired or replaced free of charge. | Machines to be free of any defects caused by materials and workmanship | All Parts and Labour | All Parts and Labour | All Parts and Labour | All Parts and Labour | All Parts and Labour | Free of charge labour and parts
Please state any conditions of the warranty | • Damage resulting from transportation, fire, accident, abuse, misuse, improper use, neglect, or act of God. • Damage resulting from immersion in or exposure to chemicals, liquids or dirt, extremes of climate, fungus or excessive wear and tear. | Standard Terms and conditions | See above | See above | Warranty period: duration The GAMA initial warranty period shall follow local minimum statutory requirements (E.g. one year in the UK). A yearly extended warranty cover is to be offered as a GAMA product, applicable from the end of the initial statutory period onwards. Warranty claim exclusions The following events will void claims during the warranty period: • Equipment misuse • Neglect or negligence • Accidental damage • Improper storage • Unauthorised alteration • Historical maintenance schedule in breach of GAMA’s guidelines • Force majeure events Service level agreement outline • Yearly preventative maintenance service included at the end each 12-month period • In the event of a fault, site visit within 48 hours from callout for on-site repairs • Temporary replacement unit is issued within 72 hours from callout should repairs not be viable in the first site visit • Customer equipment replaced back on-site once repairs have been completed off-site |  |  | After the frist 12 months, warranty is subject to the systems being routinely serviced by Inivos under the usual service level agreement | The Oxypharm Nocospray Appliance carries a comprehensive warranty for a period of 2 years from the original date of purchase. This warranty is subject to the appliance being used in accordance with the standard operating instructions issued by the supplier at the time of purchase and providing only Oxypharm chemicals have been used. Any fault or defective parts that have been caused by, or are a direct result of, operator misuse invalidates this warranty. The Nocospray appliance must only be repaired, serviced and/or calibrated by Oxypharm trained engineers using genuine Oxypharm parts. The removal of the Nocospray base plate and/or the use of non genuine Oxypharm parts invalidates this warranty. Upon making a claim under this warranty, the customer is obliged to inform their supplier of the nature and reason for claim plus provide the appliance serial number which is located on the base plate of the appliance. The Supplier or Oxypharm trained Engineer will advise the customer if the appliance will be repaired or replaced. The second of the 2 year period of warranty will only be effective if the customer ensures each appliance has received a Calibration and Service by trained Oxypharm engineers on or around the end of the first year. | The Oxypharm Nocospray Appliance carries a comprehensive warranty for a period of 2 years from the original date of purchase. This warranty is subject to the appliance being used in accordance with the standard operating instructions issued by the supplier at the time of purchase and providing only Oxypharm chemicals have been used. Any fault or defective parts that have been caused by, or are a direct result of, operator misuse invalidates this warranty. The Nocospray appliance must only be repaired, serviced and/or calibrated by Oxypharm trained engineers using genuine Oxypharm parts. The removal of the Nocospray base plate and/or the use of non genuine Oxypharm parts invalidates this warranty. Upon making a claim under this warranty, the customer is obliged to inform their supplier of the nature and reason for claim plus provide the appliance serial number which is located on the base plate of the appliance. The Supplier or Oxypharm trained Engineer will advise the customer if the appliance will be repaired or replaced. The second of the 2 year period of warranty will only be effective if the customer ensures each appliance has received a Calibration and Service by trained Oxypharm engineers on or around the end of the first year. | Machine wear and tear is not covered | The only condition of the warranty is normal, validated use in accordance with supplier/manufacturer recommendations. | The only condition of the warranty is normal, validated use in accordance with supplier/manufacturer recommendations. | The only condition of the warranty is normal, validated use in accordance with supplier/manufacturer recommendations. | The only condition of the warranty is normal, validated use in accordance with supplier/manufacturer recommendations. | The only condition of the warranty is normal, validated use in accordance with supplier/manufacturer recommendations. | Exclusion would include user error & accidental damage
Outside the warranty period
Please provide details of major components likely to require replacement in a period of not less than 7 years from the date of handover of the installation to the Trust. | Biocide spraying system calibration, compressor’s filter, liquid circuit filter, liquid circuit pipes and air pipes. | none | Power cables, main vaporiser core, dosing pump | PCB boards on the main unit, distribution fan, main vaporiser core & dosing pump | To ensure optimal efficacy levels a yearly preventative maintenance schedule is required (standard issue with a warranty agreement) which includes full filter set replacement. | None | None | The pump is likely to require replacement within 7 years, along with the main carbon filter / deactivation unit. We recommend our service level agreements for periods outside the warranty as this offers full replacement of any faulty components without additional charge | After approximately 700 hours operational use, it is recommended the internal motor of the machine should be replaced. | After approximately 700 hours operational use, it is recommended the internal motor of the machine should be replaced. | Transducer, Touch screen & Fan | None. All components with limited life-span are replaced in a time well-before their expected failure as part of the maintenance and servicing programme | None. All components with limited life-span are replaced in a time well-before their expected failure as part of the maintenance and servicing programme | None. All components with limited life-span are replaced in a time well-before their expected failure as part of the maintenance and servicing programme | None. All components with limited life-span are replaced in a time well-before their expected failure as part of the maintenance and servicing programme | None. All components with limited life-span are replaced in a time well-before their expected failure as part of the maintenance and servicing programme | GL-323 x 2 GL-225 x 1 GL-120 x 1 GL-340/0 GL-212 x 1 GL-210/00 x 1 1000.550 (maintenance kit)
Point of Sale Maintenance
Do you offer point of sale maintenance packages? | Yes | Yes | Yes | Yes | An extended warranty agreement is available for sale and applicable from year two of ownership onwards. An extended warranty agreement carries the same benefits as the first-year statutory warranty with regards to SLA levels, number of visits, lead times etc. | Yes | Yes | Not sure what is meant by this, please confirm | Yes, Extended warranty Contracts and Annual Service Contracts are available to purchase. | Yes, Extended warranty Contracts and Annual Service Contracts are available to purchase. | YES | Not Applicable |  |  |  |  | Yes

## W8 - Reference sites
Reference Site Information
Tender Name | Environmental Decontamination
Lot Number | 1
Lot Name | Hydrogen Peroxide Vapour Disinfection Systems/Fogging Machines
Details provided in this section may be shared by NHS Supply Chain with customers requesting reference sites
Those customers who qualify for 'Reference site' pricing should be listed here
Please advise locations and year of supply, all sites must be within the last 3 years
Amity | Amity |  |  |  |  | Bioquell | Bioquell |  | Gama | Gama |  | Infinite | Infinite | Inivos | Inivos |  | Innotec | Innotec | Sanital Med | Sanital Med | UTB | UTB |  | Wassenburg Medical | Wassenburg Medical
Reference Site Details | Response |  | Reference Site Details | Response |  | Reference Site Details | Response |  | Reference Site Details | Response |  | Reference Site Details | Response | Reference Site Details | Response |  | Reference Site Details | Response | Reference Site Details | Response | Reference Site Details | Response |  | Reference Site Details | Response
• Royal Glamorgan Hospital • Princesses of Wales Hospital • Princes Charles Hospital • Ysbyty Cwm Cynon Hospital • Ysbyty Cwm Rhondda Hospital | Tender was awarded in 2017 but since then many more units have been purchased on a annual basis. In total | Reference Site 1 | Wrexham Hospital | Customer specific details avaliable upon request. | Reference Site 1 | Guys & St Thomas' NHS Foundation Trust, Westminster Bridge Rd, London SE1 7EH | Contact Name: Glawdys Chukukere BQ-50, ProteQ, Pod & RBDS users | Reference Site 1 | Scarlet HPV Machines provided during the last 3 years | Betsi Cadwaladr University Health Board | Reference Site 1 | not filled in | not filled in | Addenbrookes Hospital Cambridge University Hospitals NHS Foundation Trust | Andrew Eady - Housekeeping Manager, Medirest Andrew.Eady@compass-group.co.uk 01223 348473 | Reference Site 1 | University Hospitals of Derby and Burton NHS Foundation Trust | Ongoing supply of Oxy’Pharm’s Nocotech airborne decontamination system - devices and consumables. Commenced supply July 2012. Supply has been continual and is current. | not filled in | not filled in | OXFORD UNIVERSITY . Including The Lister Institute, developers of the SARS-CoV-2, vaccine and other Departments. |  | Reference Site 1 | Royal Free London NHS Trust |  | Reference Site 1
|  |  | South Wales NHS | Customer specific details avaliable upon request. |  |  |  |  | Scarlet HPV Machines provided during the last 3 years | Southport & Ormskirk Nhs Hospital Trust | [10 empty] | Biological Safety Cabinet and Laboratory Decontamination - 2019
Queen Victoria Hospital | Purchased in 2021 | Reference Site 2 |  |  | Reference Site 2 | The London Clinic Hospital, 20 Devonshire Pl, London W1G 6BW | Contact Name: Blessed Quansah BQ-50 / Proactive RBDS Service users | Reference Site 2 | Scarlet HPV Machines provided during the last 3 years | Epsom and St Helier University Hospitals | Reference Site 2 |  |  | Bradford Teaching Hospitals NHS Foundation Trust | Rachel Sorhaindo - Facilities Manager rachel.sorhaindo@bthft.nhs.uk 07947 744926 | Reference Site 2 | Wirral University Teaching Hospital NHS Foundation Trust | Ongoing supply of Oxy’Pharm’s Nocotech airborne decontamination system - devices and consumables. Commenced supply March 2014. Supply has been continual and is current. |  |  | SynGenta UK Limited, International Research Centre, March 2020 |  | Reference Site 2 | Intergrated Facilities Management Bolton Ltd |  | Reference Site 2
[9 empty] | Scarlet HPV Machines provided during the last 3 years | Princess Alexandra Hospital NHS Trust | [10 empty] | Multiple laboratory decontamination to > 6Log reduction
[17 empty] | ISS Mediclean | Ongoing supply of Oxy’Pharm’s Nocotech airborne decontamination system - devices and consumables. Commenced supply February 2018. Supply has been continual and is current.
Please insert more sites as required | [13 empty] | Stockport NHS Foundation Trust | Joanne Morris - Support Services Manager joanne.morris@stockport.nhs.uk 0161 419 5167 | Reference Site 3 |  | Sites supplied:- Royal Liverpool Hospital, Clatterbridge Cancer Centre, Walton Centre Hospital, Broadgreen Hospital, Macclesfield General Hospital. |  |  | UNIVERSITY of GLASGOW - Virus Research Centre June 2019 |  | Reference Site 2
Multiple rooms and IVC unit with 2 rack of 60 cages to a >6Log reduction
CHU of Grenoble
Deploying Multiple Phileas Units for Rooms previously occupied by HCAI patients
HOPITAL, Paris
Deploying Multiple Phileas Units for Rooms previously occupied by HCAI patients